POLYTECHNIC OF TURIN A.Y2021/22

Course on Designing Biomedical Programmable Devices

Contact temperature gauge

GROUP 17

Scorrano Luca 296047

Sorrentino Federica 238978

SUMMARY

- Project Specifications	3

  - Enhancement Options	3

- Hardware and software component block diagram 4

  - Software block diagram	4

  - Hardware block diagram	5

- Circuit diagram	5

- Sensor Validation 7

- Software description	9

  - Introduction	9

5.3 Flow-chart	11

- Specification check and conclusions	14

Appendix	17

# Project Specifications

The following project requires the realisation of a temperature meter capable of providing the skin temperature when the LM35 sensor is placed in contact with the skin. It must meet the following specifications:

  - Supply voltage: 5V

  - Check the supply voltage at approximately 1s intervals. Should it fall below a certain programmed threshold value, generate a continuous alarm signal with LED:

Val > 4.2V           LED off

Val < 4.2V           LED on

The status of the LED should indicate the range in which the supply voltage is.

  - Temperature measurement (in degrees Celsius and Fahrenheit) at 500 ms intervals, with

resolution of 0.1 °C and 0.1 °F.

  - Reading the measurement on three seven-segment displays, providing the result in the 30°C and 45°C range.

## Enhancement Options

    - By using a mechanical switch, one wants to be able to switch the display of the measurement result from Celsius to Fahrenheit and vice versa.

    - Verification of supply voltage with double threshold alarm system:

Val > 4.8V            LED off

4.8V < Val < 4.2V           LED access in flashing mode

Val < 4.2V           LED on

# Hardware and software component block diagram

## Software block diagram

## Hardware block diagram

Trimmer

Trimmer

Micro

Button

Micro

Button

# Electrical diagram

The circuit was realised on two breadboards of 400 and 830 points.

The operation of the temperature meter is based on the integrated LM35DZ temperature sensor, which is appropriately connected to the ARDUINO UNO REV.3 board. It has an ATmega328P microcontroller inside, which is implemented in such a way that it can process the temperature information, expressed in voltage.

The LM35DZ sensor data sheet shows the following characteristics:

- Current consumption 60µA plus output current

- Maximum output current 10mA

- Output voltage 10mV/°C

- Measuring range -55°C up to 150°C

A 100nF bypass capacitor was inserted between the power supply pin and the ground reference of each CD4511BE integrated circuit. This was done in order to guarantee sufficient working voltage to the integrated circuit and to avoid, when the switch is switched, a zero supply due to the instantaneous increase of current in the branch.

In order to maximise the dynamics of the ADC, taking into account the input voltage values received from the sensor (order of hundreds of mV) and to obtain a resolution of 0.1 °C, during the temperature measurement and reading phase, the ADC was set to deliver a supply voltage of 1.1 V and used 10 bits for the conversion of the result.

𝑅𝐴/𝐷 =

1.1 𝑉

210

# Sensor Validation

  - LED alarm activation check: in order to vary the supply voltage supplied by the ADC, an alarm circuit including a potentiometer trimmer has been fitted. Acting on the latter's shaft, the circuit varies the voltage between the slider, connected to a channel of the ADC, and the reference between Vdc and GND, attesting to the LED's switching on below the set threshold.

  - Functional verification of the sensor: in order to ascertain the realistic functioning of the thermometer, a calibration by comparison was chosen. A digital electric household thermometer was used as a reference instrument. Both sensors were placed in contact with the wall of a glass of hot water and, taking advantage of the natural transfer of heat from the liquid to the external environment which causes a progressive decrease in temperature, a table was constructed with 5 measurements comparing, respectively, the reference value and the measured sensor value. (see Figure 1 and Table 1)

Table 1

| Reference value (°C) | Sensor value (°C) |
| --- | --- |
| 42 | 44 |
| 39 | 41 |
| 35,6 | 37,1 |
| 35 | 36,4 |
| 33,6 | 35,9 |

Calibration for comparison

50

45

40

35

30

25

20

15

10

5

0

1

2

3

4

5

Reference value(°C)

Sensor value (°C)

Calibration for comparison

50

45

40

35

30

25

20

15

10

5

0

1

2

3

4

5

Reference value(°C)

Sensor value (°C)

2

2

2

2

2.3

2.3

Figure 1: Calibration by comparison

  - Verification of correct operation of the button: by physically changing the state of the button, the changeover from Celsius to Fahrenheit was verified correctly

# Software description

## Introduction

To programme the microcontroller and drive the entire circuit, the Atmel Studio 7.0 development environment was used. It allows intervention at the level of the individual registers and peripherals of the microcontroller. The execution software was realised in Assembler language and subsequently loaded into the Flash Program Memory of the ATmega328P using the ARDUINO UNO REV.3 board.

It is present in the development environment:

    - an editor, which allows you to write code

    - an assembler, which makes it possible to switch between Assembler and machine language programmes

    - a simulator, which makes it possible to simulate what is happening in the microC and verify its actual operation.

  - Block diagram

## Flow-chart

Flowchart routine measure

MEASURE

ADSC=0

YES

NO

Check_conv

VERIFICATION OF END OF ADC CONVERSION

START OF ADC

CONVERSION

START OF ADC

CONVERSION

END OF CONVERSION

END OF CONVERSION

I READ THE OUTPUT VALUE IN 4-BIT STEPS AND JUSTIFY TO THE LEFT

display

display

I SHOW THE RESULT ON THE DISPLAY

I SHOW THE RESULT ON THE DISPLAY

Interrupt Response Subroutines

mp  STACK

mp  STACK

SREG STACK

INITIALISE TCNT0

INITIALISE TCNT0

DEC VAR_CHECK DEC VAR_CHECK2

DEC VAR_CHECK DEC VAR_CHECK2

STACK  SREG

STACK  SREG

reti

STACK  mp

# Specific verification and conclusions

The temperature meter, following the various tests to which it has been subjected, is found to be functional and adheres to all design specifications, with the exception of the complete conversion to degrees Farenahit. In fact, having only three displays, it was decided during the design phase to meet the required specification of 0.1 F resolution and therefore not to show the figure corresponding to hundreds of degrees Farenhait, considering it to be below 99.9 F onwards. To remedy this, an extra 7-segment display could be used.

# Appendix: Implemented Assembly Code

; contact_temperature_meter.asm

;

; Definition of registers used

.DEF	mp 	= R16		; working register

.	DEFmp1 		= R17 ; secondary work register (generic)

.DEF	var_check 	= R18		; var_check is the variable decremented by the interrupt response subroutine for checking the supply voltage

.	DEFvar_check2 	= R19		; var_check2 is the variable decremented by the interrupt response subroutine for temperature control

.DEFvar_check3 	= R25		; var_check3 is the variable that counts the anti-rebound of the button

.	DEFflags		= R20		; flag ultilised for table change

.DEF	flag_button = R24		; flag used to keep track of when the button was pressed

.	DEFdly1		= R21		; varibles used to implement the delay for the voltage change from 5V to 1.1 V between channel ADC0 and channel ADC1

.	DEFdly2		= R22

.	DEFdly3		= R23

.EQU TSH_BAT		= 210		; threshold value set for supply voltage control

.	EQUT_CHECK 	= 100		; constant defining the check range of the supply voltage

.	EQUT_CHECK2 	= 50		; constant defining the voltage sampling interval in multiples of 10ms

.	EQUTABLEN		= 1024		; table length expressed in groups of 4 bytes (not used in this version)

;

;

.CSEG

.ORG 	0x1FFF				; defines the beginning of the table to be written in flash

;

;LUT IN GRADES

tableC:

.	db1,	0,0

.	db2,	0,0

.	db3,	0,0

.	db4,	0,0

.	db5,	0,0

.	db6,	0,0

.	db7,	0,0

.	db8,	0,0

.	db9,	0,0

.	db10,	0,0

.	db11,	0,0

.	db12,	0,0

.	db13,	0,0

.	db14,	0,0

.	db15,	0,0

.	db16,	0,0

.	db17,	0,0

.	db18,	0,0

.	db19,	0,0

.	db20,	0,0

.	db21,	0,0

.	db22,	0,0

.	db23,	0,0

.	db24,	0,0

.	db25,	0,0

.	db26,	0,0

.	db27,	0,0

.	db28,	0,0

.	db29,	0,0

.	db30,	0,0

.	db31,	0,0

.	db32,	0,0

.	db33,	0,0

.	db34,	0,0

.	db35,	0,0

.	db36,	0,0

.	db37,	0,0

.	db38,	0,0

.	db39,	0,0

.	db40,	0,0

.	db41,	0,0

.	db42,	0,0

.	db43,	0,0

.	db44,	0,0

.	db45,	0,0

.	db46,	0,0

.	db47,	0,0

.	db48,	0,0

.	db49,	0,0

.	db50,	0,0

.	db51,	0,0

.	db52,	0,0

.	db53,	0,0

.	db54,	0,0

.	db55,	0,0

.	db56,	0,0

.	db57,	0,0

.	db58,	0,0

.	db59,	0,0

.	db60,	0,0

.	db61,	0,0

.	db62,	0,0

.	db63,	0,0

.	db64,	0,0

.	db65,	0,0

.	db66,	0,0

.	db67,	0,0

.	db68,0	,0

.	db69,	0,0

.	db70,	0,0

.	db71,	0,0

.	db72,	0,0

.	db73,	0,0

.	db74,	0,0

.	db75,	0,0

.	db76,	0,0

.	db77,	0,0

.	db78,	0,0

.	db79,	0,0

.	db80,	0,0

.	db81,	0,0

.	db82,	0,0

.	db83,	0,0

.	db84,	0,0

.	db85,	0,0

.	db86,	0,0

.	db87,	0,0

.	db88,	0,0

.	db89,	0,0

.	db90,	0,0

.	db91,	0,0

.	db92,	0,0

.	db93,0	,0

.	db94,	0,0

.	db95,	0,0

.	db96,	0,0

.	db97,	0,0

.	db98,	0,0

.	db99,	0,0

.	db100,	0,0

.	db101,	0,0

.	db102,	0,0

.	db103,	0,0

.	db104,	0,0

.	db105,	0,0

.	db106,	0,0

.	db107,	0,0

.	db108,	0,0

.	db109,	0,0

.	db110,	0,0

.	db111,	0,0

.	db112,	0,0

.	db113,0,0

.	db114,	0,0

.	db115,	0,0

.	db116,	0,0

.	db117,	0,0

.	db118,	0,0

.	db119,	0,0

.	db120,	0,0

.	db121,	0,0

.	db122,	0,0

.	db123,0,0

.	db124,	0,0

.	db125,0,0

.	db126,	0,0

.	db127,	0,0

.	db128,0,0

.	db129,	0,0

.	db130,	0,0

.	db131,	0,0

.	db132,	0,0

.	db133,	0,0

.	db134,	0,0

.	db135,	0,0

.	db136,	0,0

.	db137,	0,0

.	db138,	0,0

.	db139,	0,0

.	db140,	0,0

.	db141,	0,0

.	db142,	0,0

.	db143,	0,0

.	db144,	0,0

.	db145,	0,0

.	db146,	0,0

.	db147,	0,0

.	db148,	0,0

.	db149,	0,0

.	db150,	0,0

.	db151,	0,0

.	db152,	0,0

.	db153,	0,0

.	db154,	0,0

.	db155,	0,0

.	db156,	0,0

.	db157,	0,0

.	db158,	0,0

.	db159,	0,0

.	db160,	0,0

.	db161,	0,0

.	db162,	0,0

.	db163,	0,0

.	db164,	0,0

.	db165,	0,0

.	db166,	0,0

.	db167,	0,0

.	db168,	0,0

.	db169,	0,0

.	db170,	0,0

.	db171,	0,0

.	db172,	0,0

.	db173,	0,0

.	db174,	0,0

.	db175,	0,0

.	db176,	0,0

.	db177,	0,0

.	db178,	0,0

.	db179,	0,0

.	db180,	0,0

.	db181,	0,0

.	db182,	0,0

.	db183,	0,0

.	db184,	0,0

.	db185,	0,0

.	db186,	0,0

.	db187,	0,0

.	db188,	0,0

.	db189,	0,0

.	db190,	0,0

.	db191,	0,0

.	db192,	0,0

.	db193,	0,0

.	db194,	0,0

.	db195,	0,0

.	db196	,3,0,0

.	db197	,3,0,0

.	db198,	3,0,0

.	db199,	3,0,0

.	db200,	3,0,0

.	db201,	3,0,0

.	db202,	3,0,0

.	db203,	3,0,0

.	db204,	3,0,0

.	db205,	3,0,0

.	db206	,3,0,0

.	db207	,3,0,0

.	db208	,3,0,0

.	db209,	3,0,0

.	db210,	3,0,0

.	db211,	3,0,0

.	db212	,3,0,0

.	db213	,3,0,0

.	db214,	3,0,0

.	db215,	3,0,0

.	db216,	3,0,0

.	db217	,3,0,0

.	db218,	3,0,0

.	db219,	3,0,0

.	db220,	3,0,0

.	db221,	3,0,0

.	db222,	3,0,0

.	db223,	3,0,0

.	db224,	3,0,0

.	db225	,3,0,0

.	db226,	3,0,0

.	db227	,3,0,0

.	db228,	3,0,0

.	db229,	3,0,0

.	db230,	3,0,0

.	db231,	3,0,0

.	db232	,3,0,0

.	db233	,3,0,0

.	db234	,3,0,0

.	db235	,3,0,0

.	db236,	3,0,0

.	db237,	3,0,0

.	db238,	3,0,0

.	db239,	3,0,0

.	db240,	3,0,0

.	db241,	3,0,0

.	db242	,3,0,0

.	db243,	3,0,0

.	db244,	3,0,0

.	db245	,3,0,0

.	db246	,3,0,0

.	db247	,3,0,0

.	db248	,3,0,0

.	db249,	3,0,0

.	db250,	3,0,0

.	db251,	3,0,0

.	db252,	3,0,0

.	db253,	3,0,0

.	db254,	3,0,0

.	db255,	3,0,0

.	db256,	3,0,0

.	db257,	3,0,0

.	db258,	3,0,0

.	db259,	3,0,0

.	db260	,3,0,0

.	db261,	3,0,0

.	db262	,3,0,0

.	db263,	3,0,0

.	db264	,3,0,0

.	db265	,3,0,0

.	db266	,3,0,0

.	db267,	3,0,0

.	db268	,3,0,0

.	db269,	3,0,0

.	db270	,3,0,0

.	db271,	3,0,0

.	db272	,3,0,0

.	db273,	3,0,0

.	db274,	3,0,0

.	db275	,3,0,0

.	db276,	3,0,0

.	db277	,3,0,0

.	db278	,3,0,0

.	db279,	3,0,0

.	db280,	3,0,0

.	db281,	3,0,0

.	db282,	3,0,0

.	db283,	3,0,0

.	db284,	3,0,0

.	db285	,3,0,0

.	db286,	3,0,0

.	db287	,3,0,0

.	db288	,3,0,0

.	db289,	3,0,0

.	db290,	3,0,0

.	db291,	3,0,0

.	db292	,3,0,0

.	db293,	3,0,0

.	db294	,3,0,0

.	db295	,3,0,0

.	db296	,3,0,0

.	db297,	3,0,0

.	db298	,3,0,0

.	db299	,3,0,0

.	db300,	3,0,0

.	db301,	3,0,1

.	db302	,3,0,2

.	db303,	3,0,3

.	db304	,3,0,4

.	db305,	3,0,5

.	db306,	3,0,6

.	db307	,3,0,7

.	db308	,3,0,8

.	db309,	3,0,9

.	db310,	3,1,0

.	db311,	3,1,0

.	db312	,3,1,1

.	db313	,3,1,2

.	db314,	3,1,3

.	db315,	3,1,4

.	db316,	3,1,5

.	db317	,3,1,6

.	db318	,3,1,7

.	db319	,3,1,8

.	db320	,3,1,9

.	db321,	3,2,0

.	db322	,3,2,1

.	db323,	3,2

.	db324,	3,2,3

.	db325	,3,2,4

.	db326,	3,2,5

.	db327,	3,2,6

.	db328	,3,2,7

.	db329,	3,2,8

.	db330	,3,2,9

.	db331,	3,0

.	db332,	3,3,1

.	db333,	3,3,2

.	db334,	3,3

.	db335,	3,4

.	db336,	3,5

.	db337,	3,3,6

.	db338,	3,3,7

.	db339,	3,3,8

.	db340,	3,3,9

.	db341,	3,4,0

.	db342,	3,4,1

.	db343,	3,4,2

.	db344,	3,4,3

.	db345	,3,4

.	db346	,3,4,5

.	db347	,3,4,6

.	db348	,3,4,7

.	db349,	3,4,8

.	db350	,3,4,9

.	db351,	3,5,0

.	db352,	3,5,1

.	db353,	3,5,2

.	db354,	3,5,3

.	db355,	3,5,4

.	db356,	3,5

.	db357	,3,5,6

.	db358	,3,5,7

.	db359,	3,5,8

.	db360	,3,5,9

.	db361,	3,6,0

.	db362	,3,6,1

.	db363	,3,6,2

.	db364,	3,6,3

.	db365	,3,6,4

.	db366	,3,6,5

.	db367	,3,6

.	db368	,3,6,7

.	db369,	3,6,8

.	db370	,3,6,9

.	db371,	3,7,0

.	db372	,3,7,1

.	db373	,3,7,2

.	db374,	3,7,3

.	db375	,3,7,4

.	db376,	3,7,5

.	db377	,3,7,6

.	db378	,3,7

.	db379	,3,7,8

.	db380	,3,7,9

.	db381,	3,8,0

.	db382,	3,8,1

.	db383	,3,8,2

.	db384,	3,8,3

.	db385	,3,8,4

.	db386,	3,8,5

.	db387	,3,8,6

.	db388	,3,8,7

.	db389,	3,8

.	db390	,3,8,9

.	db391,	3,9,0

.	db392,	3,9,1

.	db393,	3,9,2

.	db394,	3,9,3

.	db395	,3,9,4

.	db396,	3,9,5

.	db397,	3,9,6

.	db398,	3,9,7

.	db399,	3,9,8

.	db400	,3,9

.	db401,	4,0,0

.	db402,	4,0,1

.	db403,	4,0,2

.	db404,	4,0,3

.	db405,	4,0,4

.	db406,	4,0,5

.	db407,	4,0,6

.	db408,	4,0,7

.	db409,	4,0,8

.	db410,	4,0,9

.	db411,	4,1,0

.	db412,	4,1,1

.	db413	,4,1,2

.	db414,	4,1,3

.	db415,	4,1,4

.	db416,	4,1,5

.	db417	,4,1,6

.	db418,	4,1,7

.	db419,	4,1,8

.	db420	,4,1,9

.	db421,	4,2,0

.	db422,	4,2,1

.	db423,	4,2,2

.	db424	,4,2,3

.	db425,	4,2,4

.	db426,	4,2,5

.	db427,	4,2,6

.	db428,	4,2,7

.	db429,	4,2,8

.	db430,	4,2,9

.	db431,	4,3,0

.	db432	,4,3,1

.	db433,	4,3,2

.	db434	,4,3,3

.	db435,	4,3,4

.	db436,	4,3,5

.	db437,	4,3,6

.	db438,	4,3,7

.	db439,	4,3,8

.	db440,	4,3,9

.	db441,	4,0

.	db442,	4,1

.	db443,	4,4,2

.	db444,	4,3

.	db445,	4,4

.	db446,	4,5

.	db447,	4,4,6

.	db448,	4,4,7

.	db449,	4,4,8

.	db450,	4,4,9

.	db451,	4,5,0

.	db452,	0,0

.	db453,	0,0

.	db454,	0,0

.	db455,	0,0

.	db456,	0,0

.	db457,	0,0

.	db458,	0,0

.	db459,	0,0

.	db460,	0,0

.	db461,	0,0

.	db462,	0,0

.	db463,	0,0

.	db464,	0,0

.	db465,	0,0

.	db466,	0,0

.	db467,	0,0

.	db468,	0,0

.	db469,	0,0

.	db470,	0,0

.	db471,	0,0

.	db472,	0,0

.	db473,	0,0

.	db474,	0,0

.	db475,	0,0

.	db476,	0,0

.	db477,	0,0

.	db478,	0,0

.	db479,	0,0

.	db480,	0,0

.	db481,	0,0

.	db482,	0,0

.	db483,	0,0

.	db484,	0,0

.	db485,	0,0

.	db486,	0,0

.	db487,	0,0

.	db488,	0,0

.	db489,	0,0

.	db490,	0,0

.	db491,	0,0

.	db492,	0,0

.	db493,	0,0

.	db494,	0,0

.	db495,	0,0

.	db496,	0,0

.	db497,	0,0

.	db498,	0,0

.	db499,	0,0

.	db500,	0,0

.	db501,	0,0

.	db502,	0,0

.	db503,	0,0

.	db504,	0,0

.	db505,	0,0

.	db506,	0,0

.	db507,	0,0

.	db508,	0,0

.	db509,	0,0

.	db510,	0,0

.	db511,	0,0

.	db512,	0,0

.	db513,	0,0

.	db514,	0,0

.	db515,	0,0

.	db516,	0,0

.	db517,	0,0

.	db518,	0,0

.	db519,	0,0

.	db520,	0,0

.	db521,	0,0

.	db522,	0,0

.	db523,	0,0

.	db524,	0,0

.	db525,	0,0

.	db526,	0,0

.	db527,	0,0

.	db528,	0,0

.	db529,	0,0

.	db530,	0,0

.	db531,	0,0

.	db532,	0,0

.	db533,	0,0

.	db534,	0,0

.	db535,	0,0

.	db536,	0,0

.	db537,	0,0

.	db538,	0,0

.	db539,	0,0

.	db540,	0,0

.	db541,	0,0

.	db542,	0,0

.	db543,	0,0

.	db544,	0,0

.	db545,	0,0

.	db546,	0,0

.	db547,	0,0

.	db548,	0,0

.	db549,	0,0

.	db550,	0,0

.	db551,	0,0

.	db552,	0,0

.	db553,	0,0

.	db554,	0,0

.	db555,	0,0

.	db556,	0,0

.	db557,	0,0

.	db558,	0,0

.	db559,	0,0

.	db560,	0,0

.	db561,	0,0

.	db562,	0,0

.	db563,	0,0

.	db564,	0,0

.	db565,	0,0

.	db566,	0,0

.	db567,	0,0

.	db568,	0,0

.	db569,	0,0

.	db570,	0,0

.	db571,	0,0

.	db572,	0,0

.	db573,	0,0

.	db574,	0,0

.	db575,	0,0

.	db576,	0,0

.	db577,	0,0

.	db578,	0,0

.	db579,	0,0

.	db580,	0,0

.	db581,	0,0

.	db582,	0,0

.	db583,	0,0

.	db584,	0,0

.	db585,	0,0

.	db586,	0,0

.	db587,	0,0

.	db588,	0,0

.	db589,	0,0

.	db590,	0,0

.	db591,	0,0

.	db592,	0,0

.	db593,	0,0

.	db594,	0,0

.	db595,	0,0

.	db596,	0,0

.	db597,	0,0

.	db598,	0,0

.	db599,	0,0

.	db600	,0,0

.	db601,	0,0

.	db602,	0,0

.	db603,	0,0

.	db604,	0,0

.	db605,	0,0

.	db606,	0,0

.	db607,	0,0

.	db608,	0,0

.	db609,	0,0

.	db610,	0,0

.	db611,	0,0

.	db612,	0,0

.	db613,	0,0

.	db614,	0,0

.	db615,	0,0

.	db616,	0,0

.	db617,	0,0

.	db618,	0,0

.	db619,	0,0

.	db620,	0,0

.	db621,	0,0

.	db622,	0,0

.	db623,	0,0

.	db624,	0,0

.	db625,	0,0

.	db626,	0,0

.	db627,	0,0

.	db628,	0,0

.	db629,	0,0

.	db630,	0,0

.	db631,	0,0

.	db632,	0,0

.	db633,	0,0

.	db634	,0,0

.	db635,	0,0

.	db636,	0,0

.	db637,	0,0

.	db638,	0,0

.	db639,	0,0

.	db640,	0,0

.	db641,	0,0

.	db642,	0,0

.	db643,	0,0

.	db644,	0,0

.	db645,	0,0

.	db646,	0,0

.	db647,	0,0

.	db648,	0,0

.	db649,	0,0

.	db650,	0,0

.	db651,	0,0

.	db652,	0,0

.	db653,	0,0

.	db654,	0,0

.	db655,	0,0

.	db656,	0,0

.	db657,	0,0

.	db658,	0,0

.	db659,	0,0

.	db660,	0,0

.	db661,	0,0

.	db662,	0,0

.	db663,	0,0

.	db664,	0,0

.	db665,	0,0

.	db666,	0,0

.	db667,	0,0

.	db668,	0,0

.	db669,	0,0

.	db670,	0,0

.	db671,	0,0

.	db672,	0,0

.	db673,	0,0

.	db674,	0,0

.	db675,	0,0

.	db676,	0,0

.	db677,	0,0

.	db678,	0,0

.	db679,	0,0

.	db680,	0,0

.	db681,	0,0

.	db682,	0,0

.	db683,	0,0

.	db684,	0,0

.	db685,	0,0

.	db686,	0,0

.	db687,	0,0

.	db688,	0,0

.	db689,	0,0

.	db690,	0,0

.	db691,	0,0

.	db692,	0,0

.	db693,	0,0

.	db694,	0,0

.	db695,	0,0

.	db696,	0,0

.	db697,	0,0

.	db698,	0,0

.	db699,	0,0

.	db700,	0,0

.	db701,	0,0

.	db702,	0,0

.	db703,	0,0

.	db704,	0,0

.	db705,	0,0

.	db706,	0,0

.	db707,	0,0

.	db708,	0,0

.	db709,	0,0

.	db710,	0,0

.	db711,	0,0

.	db712,	0,0

.	db713,0,0

.	db714,	0,0

.	db715,	0,0

.	db716,	0,0

.	db717,	0,0

.	db718,	0,0

.	db719,	0,0

.	db720,	0,0

.	db721,	0,0

.	db722,	0,0

.	db723,	0,0

.	db724, 0,0

.	db725,	0,0

.	db726,	0,0

.	db727,	0,0

.	db728,	0,0

.	db729,	0,0

.	db730,	0,0

.	db731,	0,0

.	db732,	0,0

.	db733,	0,0

.	db734, 0,0

.	db735,	0,0

.	db736,	0,0

.	db737,	0,0

.	db738,	0,0

.	db739,	0,0

.	db740,	0,0

.	db741,	0,0

.	db742,	0,0

.	db743,	0,0

.	db744,	0,0

.	db745, 0,0

.	db746,	0,0

.	db747,	0,0

.	db748,	0,0

.	db749,	0,0

.	db750,	0,0

.	db751,	0,0

.	db752,	0,0

.	db753,	0,0

.	db754,	0,0

.	db755,	0,0

.	db756,	0,0

.	db757,	0,0

.	db758,	0,0

.	db759,	0,0

.	db760,	0,0

.	db761,	0,0

.	db762,	0,0

.	db763,	0,0

.	db764, 0,0

.	db765,	0,0

.	db766, 0,0

.	db767, 0,0

.	db768, 0,0

.	db769,	0,0

.	db770,	0,0

.	db771,	0,0

.	db772,	0,0

.	db773,	0,0

.	db774,	0,0

.	db775	,0,0

.	db776,	0,0

.	db777,	0,0

.	db778, 0,0

.	db779,	0,0

.	db780,	0,0

.	db781,	0,0

.	db782,	0,0

.	db783,	0,0

.	db784,	0,0

.	db785,	0,0

.	db786,	0,0

.	db787	,0,0

.	db788,	0,0

.	db789,	0,0

.	db790,	0,0

.	db791,	0,0

.	db792,	0,0

.	db793, 0,0

.	db794,	0,0

.	db795,	0,0

.	db796,	0,0

.	db797,	0,0

.	db798,	0,0

.	db799,	0,0

.	db800,	0,0

.	db801,	0,0

.	db802,	0,0

.	db803,	0,0

.	db804,	0,0

.	db805,	0,0

.	db806,	0,0

.	db807,	0,0

.	db808,	0,0

.	db809,	0,0

.	db810,	0,0

.	db811,	0,0

.	db812,	0,0

.	db813, 0,0

.	db814,	0,0

.	db815,	0,0

.	db816,	0,0

.	db817,	0,0

.	db818,	0,0

.	db819,	0,0

.	db820,	0,0

.	db821,	0,0

.	db822,	0,0

.	db823,	0,0

.	db824, 0,0

.	db825,	0,0

.	db826, 0,0

.	db827, 0,0

.	db828,	0,0

.	db829,	0,0

.	db830,	0,0

.	db831,	0,0

.	db832,	0,0

.	db833,	0,0

.	db834	,0,0

.	db835,	0,0

.	db836,	0,0

.	db837,	0,0

.	db838,	0,0

.	db839,	0,0

.	db840,	0,0

.	db841,	0,0

.	db842,	0,0

.	db843,	0,0

.	db844,	0,0

.	db845, 0,0

.	db846,	0,0

.	db847,	0,0

.	db848,	0,0

.	db849,	0,0

.	db850,	0,0

.	db851,	0,0

.	db852,	0,0

.	db853,	0,0

.	db854,	0,0

.	db855,	0,0

.	db856,	0,0

.	db857,	0,0

.	db858, 0,0

.	db859,	0,0

.	db860,	0,0

.	db861,	0,0

.	db862,	0,0

.	db863,	0,0

.	db864, 0,0

.	db865,	0,0

.	db866,	0,0

.	db867,	0,0

.	db868, 0,0

.	db869,	0,0

.	db870,	0,0

.	db871,	0,0

.	db872,	0,0

.	db873,	0,0

.	db874,	0,0

.	db875,	0,0

.	db876,	0,0

.	db877,	0,0

.	db878, 0,0

.	db879,	0,0

.	db880,	0,0

.	db881,	0,0

.	db882,	0,0

.	db883,	0,0

.	db884,	0,0

.	db885,	0,0

.	db886,	0,0

.	db887,	0,0

.	db888,	0,0

.	db889,	0,0

.	db890,	0,0

.	db891,	0,0

.	db892,	0,0

.	db893, 0,0

.	db894,	0,0

.	db895,	0,0

.	db896,	0,0

.	db897,	0,0

.	db898,	0,0

.	db899,	0,0

.	db900,	0,0

.	db901,	0,0

.	db902,	0,0

.	db903,	0,0

.	db904,	0,0

.	db905,	0,0

.	db906,	0,0

.	db907,	0,0

.	db908,	0,0

.	db909,	0,0

.	db910,	0,0

.	db911,	0,0

.	db912,	0,0

.	db913, 0,0

.	db914,	0,0

.	db915,	0,0

.	db916,	0,0

.	db917	,0,0

.	db918,	0,0

.	db919, 0,0

.	db920,	0,0

.	db921,	0,0

.	db922	,0,0

.	db923	,0,0

.	db924, 0,0

.	db925	,0,0

.	db926,	0,0

.	db927, 0,0

.	db928	,0,0

.	db929,	0,0

.	db930,	0,0

.	db931,	0,0

.	db932	,0,0

.	db933,	0,0

.	db934	,0,0

.	db935	,0,0

.	db936,	0,0

.	db937,	0,0

.	db938,	0,0

.	db939,	0,0

.	db940,	0,0

.	db941,	0,0

.	db942,	0,0

.	db943,	0,0

.	db944,	0,0

.	db945, 0,0

.	db946,	0,0

.	db947,	0,0

.	db948,	0,0

.	db949,	0,0

.	db950,	0,0

.	db951,	0,0

.	db952,	0,0

.	db953,	0,0

.	db954,	0,0

.	db955,	0,0

.	db956,	0,0

.	db957,	0,0

.	db958,	0,0

.	db959,	0,0

.	db960,	0,0

.	db961,	0,0

.	db962,	0,0

.	db963,	0,0

.	db964,	0,0

.	db965,	0,0

.	db966,	0,0

.	db967,	0,0

.	db968, 0,0

.	db969, 0,0

.	db970,	0,0

.	db971,	0,0

.	db972	,0,0

.	db973,	0,0

.	db974,	0,0

.	db975,	0,0

.	db976,	0,0

.	db977,	0,0

.	db978, 0,0

.	db979,	0,0

.	db980,	0,0

.	db981,	0,0

.	db982,	0,0

.	db983,	0,0

.	db984,	0,0

.	db985,	0,0

.	db986,	0,0

.	db987,	0,0

.	db988,	0,0

.	db989,	0,0

.	db990,	0,0

.	db991,	0,0

.	db992,	0,0

.	db993,	0,0

.	db994,	0,0

.	db995,	0,0

.	db996,	0,0

.	db997,	0,0

.	db998,	0,0

.	db999, 0,0

.db1000 ,	0,0

.db1001 ,	0,0

.db1002 ,	0,0

.db1003 ,	0,0

.db1004 ,	0,0

.db1005 ,	0,0

.db1006 ,	0,0

.db1007 ,	0,0

.db1008 ,	0,0

.db1009 ,	0,0

.db1010 ,	0,0

.db1011 ,	0,0

.db1012 	,0,0

.db1013 ,	0,0

.db1014 ,	0,0

.db1015 ,	0,0

.db1016 ,	0,0

.db1017 ,	0,0

.db1018 ,	0,0

.db1019 ,	0,0

.db1020 ,	0,0

.db1021 ,	0,0

.db1022 ,	0,0

.db1023 ,	0,0

.db1024 ,	0,0

;

LUT IN FAHRENHEIT

.CSEG

.ORG 0x15C

tableF:

.	db1,	0,0

.	db2,	0,0

.	db3,	0,0

.	db4,	0,0

.	db5,	0,0

.	db6,	0,0

.	db7,	0,0

.	db8,	0,0

.	db9,	0,0

.	db10,	0,0

.	db11,	0,0

.	db12,	0,0

.	db13,	0,0

.	db14,	0,0

.	db15,	0,0

.	db16,	0,0

.	db17,	0,0

.	db18,	0,0

.	db19,	0,0

.	db20,	0,0

.	db21,	0,0

.	db22,	0,0

.	db23,	0,0

.	db24,	0,0

.	db25,	0,0

.	db26,	0,0

.	db27,	0,0

.	db28,	0,0

.	db29,	0,0

.	db30,	0,0

.	db31,	0,0

.	db32,	0,0

.	db33,	0,0

.	db34,	0,0

.	db35,	0,0

.	db36,	0,0

.	db37,	0,0

.	db38,	0,0

.	db39,	0,0

.	db40,	0,0

.	db41,	0,0

.	db42,	0,0

.	db43,	0,0

.	db44,	0,0

.	db45,	0,0

.	db46,	0,0

.	db47,	0,0

.	db48,	0,0

.	db49,	0,0

.	db50,	0,0

.	db51,	0,0

.	db52,	0,0

.	db53,	0,0

.	db54,	0,0

.	db55,	0,0

.	db56,	0,0

.	db57,	0,0

.	db58,	0,0

.	db59,	0,0

.	db60,	0,0

.	db61,	0,0

.	db62,	0,0

.	db63,	0,0

.	db64,	0,0

.	db65,	0,0

.	db66,	0,0

.	db67,	0,0

.	db68,0	,0

.	db69,	0,0

.	db70,	0,0

.	db71,	0,0

.	db72,	0,0

.	db73,	0,0

.	db74,	0,0

.	db75,	0,0

.	db76,	0,0

.	db77,	0,0

.	db78,	0,0

.	db79,	0,0

.	db80,	0,0

.	db81,	0,0

.	db82,	0,0

.	db83,	0,0

.	db84,	0,0

.	db85,	0,0

.	db86,	0,0

.	db87,	0,0

.	db88,	0,0

.	db89,	0,0

.	db90,	0,0

.	db91,	0,0

.	db92,	0,0

.	db93,0	,0

.	db94,	0,0

.	db95,	0,0

.	db96,	0,0

.	db97,	0,0

.	db98,	0,0

.	db99,	0,0

.	db100,	0,0

.	db101,	0,0

.	db102,	0,0

.	db103,	0,0

.	db104,	0,0

.	db105,	0,0

.	db106,	0,0

.	db107,	0,0

.	db108,	0,0

.	db109,	0,0

.	db110,	0,0

.	db111,	0,0

.	db112,	0,0

.	db113, 0,0

.	db114,	0,0

.	db115,	0,0

.	db116,	0,0

.	db117,	0,0

.	db118,	0,0

.	db119,	0,0

.	db120,	0,0

.	db121,	0,0

.	db122,	0,0

.	db123, 0,0

.	db124,	0,0

.	db125, 0,0

.	db126,	0,0

.	db127,	0,0

.	db128, 0,0

.	db129,	0,0

.	db130,	0,0

.	db131,	0,0

.	db132,	0,0

.	db133,	0,0

.	db134,	0,0

.	db135,	0,0

.	db136,	0,0

.	db137,	0,0

.	db138,	0,0

.	db139,	0,0

.	db140,	0,0

.	db141,	0,0

.	db142,	0,0

.	db143,	0,0

.	db144,	0,0

.	db145,	0,0

.	db146,	0,0

.	db147,	0,0

.	db148,	0,0

.	db149,	0,0

.	db150,	0,0

.	db151,	0,0

.	db152,	0,0

.	db153,	0,0

.	db154,	0,0

.	db155,	0,0

.	db156,	0,0

.	db157,	0,0

.	db158,	0,0

.	db159,	0,0

.	db160,	0,0

.	db161,	0,0

.	db162,	0,0

.	db163,	0,0

.	db164,	0,0

.	db165,	0,0

.	db166,	0,0

.	db167,	0,0

.	db168,	0,0

.	db169,	0,0

.	db170,	0,0

.	db171,	0,0

.	db172,	0,0

.	db173,	0,0

.	db174,	0,0

.	db175,	0,0

.	db176,	0,0

.	db177,	0,0

.	db178,	0,0

.	db179,	0,0

.	db180,	0,0

.	db181,	0,0

.	db182,	0,0

.	db183,	0,0

.	db184,	0,0

.	db185,	0,0

.	db186,	0,0

.	db187,	0,0

.	db188,	0,0

.	db189,	0,0

.	db190,	0,0

.	db191,	0,0

.	db192,	0,0

.	db193,	0,0

.	db194,	0,0

.	db195,	0,0

.	db196,	0,0

.	db197,	0,0

.	db198,	0,0

.	db199,	0,0

.	db200,	0,0

.	db201,	0,0

.	db202,	0,0

.	db203,	0,0

.	db204,	0,0

.	db205,	0,0

.	db206,	0,0

.	db207,	0,0

.	db208,	0,0

.	db209,	0,0

.	db210,	0,0

.	db211,	0,0

.	db212,	0,0

.	db213, 0,0

.	db214,	0,0

.	db215,	0,0

.	db216,	0,0

.	db217,	0,0

.	db218,	0,0

.	db219,	0,0

.	db220,	0,0

.	db221,	0,0

.	db222,	0,0

.	db223,	0,0

.	db224,	0,0

.	db225,	0,0

.	db226,	0,0

.	db227,	0,0

.	db228,	0,0

.	db229,	0,0

.	db230,	0,0

.	db231,	0,0

.	db232	,0,0

.	db233,	0,0

.	db234	,0,0

.	db235,	0,0

.	db236,	0,0

.	db237,	0,0

.	db238,	0,0

.	db239,	0,0

.	db240,	0,0

.	db241,	0,0

.	db242,	0,0

.	db243,	0,0

.	db244,	0,0

.	db245, 0,0

.	db246,	0,0

.	db247,	0,0

.	db248,	0,0

.	db249,	0,0

.	db250,	0,0

.	db251,	0,0

.	db252,	0,0

.	db253,	0,0

.	db254,	0,0

.	db255,	0,0

.	db256,	0,0

.	db257,	0,0

.	db258,	0,0

.	db259,	0,0

.	db260,	0,0

.	db261,	0,0

.	db262,	0,0

.	db263,	0,0

.	db264,	0,0

.	db265,	0,0

.	db266,	0,0

.	db267,	0,0

.	db268, 0,0

.	db269,	0,0

.	db270,	0,0

.	db271,	0,0

.	db272	,0,0

.	db273,	0,0

.	db274,	0,0

.	db275,	0,0

.	db276,	0,0

.	db277,	0,0

.	db278,	0,0

.	db279,	0,0

.	db280,	0,0

.	db281,	0,0

.	db282,	0,0

.	db283,	0,0

.	db284,	0,0

.	db285,	0,0

.	db286,	0,0

.	db287,	0,0

.	db288,	0,0

.	db289,	0,0

.	db290,	0,0

.	db291,	0,0

.	db292,	0,0

.	db293,	0,0

.	db294,	0,0

.	db295,	0,0

.	db296,	0,0

.	db297,	0,0

.	db298,	0,0

.	db299,	0,0

.	db300,	8,6,0

.	db301,	8,6,1

.	db302	,8,6,3

.	db303,	8,6,5

.	db304	,8,6,7

.	db305,	8,6,9

.	db306,	8,7,0

.	db307,	8,7,2

.	db308,	8,7,4

.	db309,	8,7,6

.	db310,	8,7,8

.	db311	,8,7,9

.	db312,	8,1

.	db313,	8,3

.	db314,	8,5

.	db315,	8,7

.	db316,	8,8

.	db317,	8,9,0

.	db318,	8,9,2

.	db319	,8,9,2

.	db320	,8,9,4

.	db321,	8,9,6

.	db322	,8,9,7

.	db323	,8,9

.	db324,	9,0,1

.	db325	,9,0,3

.	db326,	9,0,5

.	db327,	9,0,6

.	db328,	9,0,8

.	db329,	9,1,0

.	db330	,9,1,2

.	db331,	9,1,4

.	db332	,9,1,5

.	db333	,9,1,7

.	db334,	9,1,9

.	db335	,9,2,1

.	db336	,9,2,3

.	db337	,9,2,4

.	db338	,9,2,6

.	db339,	9,2,8

.	db340,	9,3,0

.	db341,	9,3,2

.	db342,	9,3

.	db343	,9,3,5

.	db344	,9,3,7

.	db345,	9,3,9

.	db346	,9,4,1

.	db347	,9,4,2

.	db348,	9,4

.	db349,	9,4,6

.	db350	,9,4,8

.	db351,	9,5,0

.	db352,	9,5,1

.	db353,	9,5,3

.	db354,	9,5

.	db355,	9,5,7

.	db356,	9,5,9

.	db357	,9,6,0

.	db358,	9,6,2

.	db359,	9,6,4

.	db360	,9,6

.	db361,	9,6,8

.	db362,	9,6,9

.	db363	,9,7,1

.	db364	,9,7,3

.	db365	,9,7,5

.	db366	,9,7,7

.	db367	,9,7,8

.	db368	,9,8,0

.	db369,	9,8,2

.	db370	,9,8,4

.	db371,	9,8,6

.	db372	,9,8,7

.	db373,	9,8,9

.	db374,	9,1

.	db375,	9,3

.	db376,	9,5

.	db377,	9,6

.	db378,	9,8

.	db379,	0,0

.	db380,	0,0,2

.	db381,	0,0,4

.	db382,	0,0,5

.	db383,	0,0,7

.	db384	,0,0,9

.	db385	,0,1,1

.	db386	,0,1,3

.	db387	,0,1,4

.	db388	,0,1,6

.	db389,	0,1,8

.	db390,	0,2,0

.	db391,	0,2,2

.	db392,	0,2,3

.	db393	,0,2,5

.	db394,	0,2,7

.	db395	,0,2,9

.	db396,	0,3,1

.	db397	,0,3,2

.	db398	,0,3,4

.	db399,	0,3,6

.	db400	,0,3,8

.	db401,	0,4,0

.	db402,	0,4,1

.	db403,	0,4,3

.	db404,	0,4,5

.	db405,	0,4,7

.	db406,	0,4,9

.	db407,	0,5,0

.	db408,	0,5,2

.	db409,	0,5,4

.	db410,	0,5,6

.	db411,	0,5,8

.	db412	,0,5,9

.	db413	,0,6,1

.	db414,	0,6,3

.	db415,	0,6,5

.	db416,	0,6,7

.	db417	,0,6,8

.	db418,	0,7,0

.	db419	,0,7,2

.	db420	,0,7,4

.	db421,	0,7,6

.	db422	,0,7,7

.	db423,	0,7,9

.	db424	,0,8,1

.	db425	,0,8,3

.	db426,	0,8,5

.	db427,	0,8,6

.	db428,	0,8

.	db429,	0,9,0

.	db430	,0,9,2

.	db431	,0,9,4

.	db432	,0,9,5

.	db433	,0,9,7

.	db434	,0,9,9

.	db435,	1,0,1

.	db436	,1,0,3

.	db437	,1,0,4

.	db438,	1,0,6

.	db439,	1,0,8

.	db440,	1,1,0

.	db441,	1,1,2

.	db442,	1,1,3

.	db443,	1,1,5

.	db444,	1,1,7

.	db445,	1,1,9

.	db446,	1,2,1

.	db447,	1,2,2

.	db448	,1,2,4

.	db449	,1,2,6

.	db450,	1,2,8

.	db451,	1,3,0

.	db452,	0,0

.	db453,	0,0

.	db454,	0,0

.	db455,	0,0

.	db456,	0,0

.	db457,	0,0

.	db458,	0,0

.	db459,	0,0

.	db460,	0,0

.	db461,	0,0

.	db462,	0,0

.	db463,	0,0

.	db464,	0,0

.	db465,	0,0

.	db466,	0,0

.	db467,	0,0

.	db468,	0,0

.	db469,	0,0

.	db470,	0,0

.	db471,	0,0

.	db472,	0,0

.	db473,	0,0

.	db474,	0,0

.	db475,	0,0

.	db476,	0,0

.	db477,	0,0

.	db478,	0,0

.	db479,	0,0

.	db480,	0,0

.	db481,	0,0

.	db482,	0,0

.	db483,	0,0

.	db484,	0,0

.	db485,	0,0

.	db486,	0,0

.	db487,	0,0

.	db488,	0,0

.	db489,	0,0

.	db490,	0,0

.	db491,	0,0

.	db492,	0,0

.	db493,	0,0

.	db494,	0,0

.	db495,	0,0

.	db496,	0,0

.	db497,	0,0

.	db498,	0,0

.	db499,	0,0

.	db500,	0,0

.	db501,	0,0

.	db502,	0,0

.	db503,	0,0

.	db504,	0,0

.	db505,	0,0

.	db506,	0,0

.	db507,	0,0

.	db508,	0,0

.	db509,	0,0

.	db510,	0,0

.	db511,	0,0

.	db512,	0,0

.	db513,	0,0

.	db514,	0,0

.	db515,	0,0

.	db516,	0,0

.	db517,	0,0

.	db518,	0,0

.	db519,	0,0

.	db520,	0,0

.	db521,	0,0

.	db522,	0,0

.	db523,	0,0

.	db524,	0,0

.	db525,	0,0

.	db526,	0,0

.	db527,	0,0

.	db528,	0,0

.	db529,	0,0

.	db530,	0,0

.	db531,	0,0

.	db532,	0,0

.	db533,	0,0

.	db534,	0,0

.	db535,	0,0

.	db536,	0,0

.	db537,	0,0

.	db538,	0,0

.	db539,	0,0

.	db540,	0,0

.	db541,	0,0

.	db542,	0,0

.	db543,	0,0

.	db544,	0,0

.	db545,	0,0

.	db546,	0,0

.	db547,	0,0

.	db548,	0,0

.	db549,	0,0

.	db550,	0,0

.	db551,	0,0

.	db552,	0,0

.	db553,	0,0

.	db554,	0,0

.	db555,	0,0

.	db556,	0,0

.	db557,	0,0

.	db558,	0,0

.	db559,	0,0

.	db560,	0,0

.	db561,	0,0

.	db562,	0,0

.	db563,	0,0

.	db564,	0,0

.	db565,	0,0

.	db566,	0,0

.	db567,	0,0

.	db568,	0,0

.	db569,	0,0

.	db570,	0,0

.	db571,	0,0

.	db572,	0,0

.	db573,	0,0

.	db574,	0,0

.	db575,	0,0

.	db576,	0,0

.	db577,	0,0

.	db578,	0,0

.	db579,	0,0

.	db580,	0,0

.	db581,	0,0

.	db582,	0,0

.	db583,	0,0

.	db584,	0,0

.	db585,	0,0

.	db586,	0,0

.	db587,	0,0

.	db588,	0,0

.	db589,	0,0

.	db590,	0,0

.	db591,	0,0

.	db592,	0,0

.	db593,	0,0

.	db594,	0,0

.	db595,	0,0

.	db596,	0,0

.	db597,	0,0

.	db598,	0,0

.	db599,	0,0

.	db600	,0,0

.	db601,	0,0

.	db602,	0,0

.	db603,	0,0

.	db604,	0,0

.	db605,	0,0

.	db606,	0,0

.	db607,	0,0

.	db608,	0,0

.	db609,	0,0

.	db610,	0,0

.	db611,	0,0

.	db612,	0,0

.	db613,	0,0

.	db614,	0,0

.	db615,	0,0

.	db616,	0,0

.	db617,	0,0

.	db618,	0,0

.	db619,	0,0

.	db620,	0,0

.	db621,	0,0

.	db622,	0,0

.	db623,	0,0

.	db624,	0,0

.	db625,	0,0

.	db626,	0,0

.	db627,	0,0

.	db628,	0,0

.	db629,	0,0

.	db630,	0,0

.	db631,	0,0

.	db632,	0,0

.	db633,	0,0

.	db634	,0,0

.	db635,	0,0

.	db636,	0,0

.	db637,	0,0

.	db638,	0,0

.	db639,	0,0

.	db640,	0,0

.	db641,	0,0

.	db642,	0,0

.	db643,	0,0

.	db644,	0,0

.	db645,	0,0

.	db646,	0,0

.	db647,	0,0

.	db648,	0,0

.	db649,	0,0

.	db650,	0,0

.	db651,	0,0

.	db652,	0,0

.	db653,	0,0

.	db654,	0,0

.	db655,	0,0

.	db656,	0,0

.	db657,	0,0

.	db658,	0,0

.	db659,	0,0

.	db660,	0,0

.	db661,	0,0

.	db662,	0,0

.	db663,	0,0

.	db664,	0,0

.	db665,	0,0

.	db666,	0,0

.	db667,	0,0

.	db668,	0,0

.	db669,	0,0

.	db670,	0,0

.	db671,	0,0

.	db672,	0,0

.	db673,	0,0

.	db674,	0,0

.	db675,	0,0

.	db676,	0,0

.	db677,	0,0

.	db678,	0,0

.	db679,	0,0

.	db680,	0,0

.	db681,	0,0

.	db682,	0,0

.	db683,	0,0

.	db684,	0,0

.	db685,	0,0

.	db686,	0,0

.	db687,	0,0

.	db688,	0,0

.	db689,	0,0

.	db690,	0,0

.	db691,	0,0

.	db692,	0,0

.	db693,	0,0

.	db694,	0,0

.	db695,	0,0

.	db696,	0,0

.	db697,	0,0

.	db698,	0,0

.	db699,	0,0

.	db700,	0,0

.	db701,	0,0

.	db702,	0,0

.	db703,	0,0

.	db704,	0,0

.	db705,	0,0

.	db706,	0,0

.	db707,	0,0

.	db708,	0,0

.	db709,	0,0

.	db710,	0,0

.	db711,	0,0

.	db712,	0,0

.	db713, 0,0

.	db714,	0,0

.	db715,	0,0

.	db716,	0,0

.	db717,	0,0

.	db718,	0,0

.	db719,	0,0

.	db720,	0,0

.	db721,	0,0

.	db722,	0,0

.	db723,	0,0

.	db724, 0,0

.	db725,	0,0

.	db726,	0,0

.	db727,	0,0

.	db728,	0,0

.	db729,	0,0

.	db730,	0,0

.	db731,	0,0

.	db732,	0,0

.	db733,	0,0

.	db734	,0,0

.	db735,	0,0

.	db736,	0,0

.	db737,	0,0

.	db738,	0,0

.	db739,	0,0

.	db740,	0,0

.	db741,	0,0

.	db742,	0,0

.	db743,	0,0

.	db744,	0,0

.	db745, 0,0

.	db746,	0,0

.	db747,	0,0

.	db748,	0,0

.	db749,	0,0

.	db750,	0,0

.	db751,	0,0

.	db752,	0,0

.	db753,	0,0

.	db754,	0,0

.	db755,	0,0

.	db756,	0,0

.	db757,	0,0

.	db758,	0,0

.	db759,	0,0

.	db760,	0,0

.	db761,	0,0

.	db762,	0,0

.	db763,	0,0

.	db764, 0,0

.	db765,	0,0

.	db766	,0,0

.	db767, 0,0

.	db768, 0,0

.	db769,	0,0

.	db770,	0,0

.	db771,	0,0

.	db772,	0,0

.	db773,	0,0

.	db774,	0,0

.	db775	,0,0

.	db776,	0,0

.	db777,	0,0

.	db778, 0,0

.	db779,	0,0

.	db780,	0,0

.	db781,	0,0

.	db782,	0,0

.	db783,	0,0

.	db784,	0,0

.	db785,	0,0

.	db786,	0,0

.	db787	,0,0

.	db788,	0,0

.	db789,	0,0

.	db790,	0,0

.	db791,	0,0

.	db792,	0,0

.	db793, 0,0

.	db794,	0,0

.	db795,	0,0

.	db796,	0,0

.	db797,	0,0

.	db798,	0,0

.	db799,	0,0

.	db800,	0,0

.	db801,	0,0

.	db802,	0,0

.	db803,	0,0

.	db804,	0,0

.	db805,	0,0

.	db806,	0,0

.	db807,	0,0

.	db808,	0,0

.	db809,	0,0

.	db810,	0,0

.	db811,	0,0

.	db812,	0,0

.	db813, 0,0

.	db814,	0,0

.	db815,	0,0

.	db816,	0,0

.	db817,	0,0

.	db818,	0,0

.	db819,	0,0

.	db820,	0,0

.	db821,	0,0

.	db822,	0,0

.	db823,	0,0

.	db824, 0,0

.	db825,	0,0

.	db826, 0,0

.	db827, 0,0

.	db828,	0,0

.	db829,	0,0

.	db830,	0,0

.	db831,	0,0

.	db832,	0,0

.	db833,	0,0

.	db834	,0,0

.	db835,	0,0

.	db836,	0,0

.	db837,	0,0

.	db838,	0,0

.	db839,	0,0

.	db840,	0,0

.	db841,	0,0

.	db842,	0,0

.	db843,	0,0

.	db844,	0,0

.	db845, 0,0

.	db846,	0,0

.	db847,	0,0

.	db848,	0,0

.	db849,	0,0

.	db850,	0,0

.	db851,	0,0

.	db852,	0,0

.	db853,	0,0

.	db854,	0,0

.	db855,	0,0

.	db856,	0,0

.	db857,	0,0

.	db858, 0,0

.	db859,	0,0

.	db860,	0,0

.	db861,	0,0

.	db862,	0,0

.	db863,	0,0

.	db864, 0,0

.	db865,	0,0

.	db866,	0,0

.	db867,	0,0

.	db868, 0,0

.	db869,	0,0

.	db870,	0,0

.	db871,	0,0

.	db872,	0,0

.	db873,	0,0

.	db874,	0,0

.	db875,	0,0

.	db876,	0,0

.	db877,	0,0

.	db878, 0,0

.	db879,	0,0

.	db880,	0,0

.	db881,	0,0

.	db882,	0,0

.	db883,	0,0

.	db884,	0,0

.	db885,	0,0

.	db886,	0,0

.	db887,	0,0

.	db888,	0,0

.	db889,	0,0

.	db890,	0,0

.	db891,	0,0

.	db892,	0,0

.	db893, 0,0

.	db894,	0,0

.	db895,	0,0

.	db896,	0,0

.	db897,	0,0

.	db898,	0,0

.	db899,	0,0

.	db900,	0,0

.	db901,	0,0

.	db902,	0,0

.	db903,	0,0

.	db904,	0,0

.	db905,	0,0

.	db906,	0,0

.	db907,	0,0

.	db908,	0,0

.	db909,	0,0

.	db910,	0,0

.	db911,	0,0

.	db912,	0,0

.	db913, 0,0

.	db914,	0,0

.	db915,	0,0

.	db916,	0,0

.	db917	,0,0

.	db918,	0,0

.	db919, 0,0

.	db920,	0,0

.	db921,	0,0

.	db922	,0,0

.	db923	,0,0

.	db924, 0,0

.	db925	,0,0

.	db926,	0,0

.	db927, 0,0

.	db928	,0,0

.	db929,	0,0

.	db930,	0,0

.	db931,	0,0

.	db932	,0,0

.	db933,	0,0

.	db934	,0,0

.	db935	,0,0

.	db936,	0,0

.	db937,	0,0

.	db938,	0,0

.	db939,	0,0

.	db940,	0,0

.	db941,	0,0

.	db942,	0,0

.	db943,	0,0

.	db944,	0,0

.	db945, 0,0

.	db946,	0,0

.	db947,	0,0

.	db948,	0,0

.	db949,	0,0

.	db950,	0,0

.	db951,	0,0

.	db952,	0,0

.	db953,	0,0

.	db954,	0,0

.	db955,	0,0

.	db956,	0,0

.	db957,	0,0

.	db958,	0,0

.	db959,	0,0

.	db960,	0,0

.	db961,	0,0

.	db962,	0,0

.	db963,	0,0

.	db964,	0,0

.	db965,	0,0

.	db966,	0,0

.	db967,	0,0

.	db968, 0,0

.	db969, 0,0

.	db970,	0,0

.	db971,	0,0

.	db972	,0,0

.	db973,	0,0

.	db974,	0,0

.	db975,	0,0

.	db976,	0,0

.	db977,	0,0

.	db978, 0,0

.	db979,	0,0

.	db980,	0,0

.	db981,	0,0

.	db982,	0,0

.	db983,	0,0

.	db984,	0,0

.	db985,	0,0

.	db986,	0,0

.	db987,	0,0

.	db988,	0,0

.	db989,	0,0

.	db990,	0,0

.	db991,	0,0

.	db992,	0,0

.	db993,	0,0

.	db994,	0,0

.	db995,	0,0

.	db996,	0,0

.	db997,	0,0

.	db998,	0,0

.	db999, 0,0

.db1000 ,	0,0

.db1001 ,	0,0

.db1002 ,	0,0

.db1003 ,	0,0

.db1004 ,	0,0

.db1005 ,	0,0

.db1006 ,	0,0

.db1007 ,	0,0

.db1008 ,	0,0

.db1009 ,	0,0

.db1010 ,	0,0

.db1011 ,	0,0

.db1012 	,0,0

.db1013 ,	0,0

.db1014 ,	0,0

.db1015 ,	0,0

.db1016 ,	0,0

.db1017 ,	0,0

.db1018 ,	0,0

.db1019 ,	0,0

.db1020 ,	0,0

.db1021 ,	0,0

.db1022 ,	0,0

.db1023 ,	0,0

.db1024 ,	0,0

;

.CSEG

.ORG 0x0000				; defines the start of the code at address 0x0000

;

; INTERRUPT VECTORS FOLLOW

;

jmp RESET 			vector 1:		Reset Handler

jmp EXT_INT0			vector 2:		IRQ0 Handler

jmp EXT_INT1			vector 3:		IRQ1 Handler

jmp PCINTR0			vector 4:		PCINT0 Handler

jmp PCINTR1			vector 5:		PCINT1 Handler

jmp PCINTR2			vector 6:		PCINT2 Handler

jmp WDT				vector 7:		Watchdog timer handler

jmp TIM2_COMPA			; vector 8:		Timer2 Compare A handler

jmp TIM2_COMPB			; vector 9:		Timer2 compare B handler

jmp TIM2_OVF			; vector 10:	Timer2 Overflow Handler

jmp TIM1_CAPT			; vector 11:	Timer1 Capture Handler

jmp TIM1_COMPA			; vector 12:	Timer1 CompareA Handler

jmp TIM1_COMPB			; vector 13:	Timer1 CompareB Handler

jmp TIM1_OVF			; vector 14:	Timer1 Overflow Handler

jmp TIM0_COMPA			; vector 15:	Timer 0 CompareA handler

jmp TIM0_COMPB			; vector 16:	Timer 0 CompareB handler

jmp TIM0_OVF 			; vector 17:	Timer0 Overflow Handler

jmp SPI_STC			; vector 18:	SPI Transfer Complete Handler

jmp USART_RXC			; vector 19:	USART RX Complete Handler

jmp USART_UDRE			; vector 20:	USART UDR Empty Handler

jmp USART_TXC			; vector 21:	USART TX Complete Handler

jmp ADC_conv			; vector 22:	ADC Conversion Complete Handler

jmp EE_RDY			; vector 23:	EEPROM Ready Handler

jmp ANA_COMP			; vector 24:	Analog Comparator Handler

jmp TWSI			vector 25:	Two-wire Serial Interface Handler

jmp SPM_RDY			; vector 26:	Store Program Memory Ready Handler

;

; END OF INTERRUPT VECTORS

;

RESET:

;

; Stack pointer initialisation to the last RAM cell

;

ldi mp,HIGH(RAMEND)

out SPH,mp

ldi mp,LOW(RAMEND)

out SPL,mp

;

; Initialisation of port B for power control, setting pin 5 as output

;

at	mp,DDRB

ldi	mp, 0b0010_0000

outDDRB	, mp

;

ldi mp, 0b1101_1111

out PORTB,mp

; Output initialisation of the first 7 bits of portD. PD0 - PD3 = ABCD; PD4 = LE hundreds; PD5 = LE ;tens; PD6 = LE units

Input initialisation of pin 7 of port D connected to the button.

ldimp	,0b0111_1111

out	DDRD,mp

;

ldi	mp, 0b0111_0000

outPORTD	, mp

;

; divides the internal clock frequency (16 MHz) by eight

;

ldi	mp, 0b1000_0000

stsCLKPR	, mp				; enables prescaler programming

ldi	mp, 0b0000_0100

stsCLKPR	, mp				; programs the division by 16 of the internal clock

;

; select prescaler step 1024 (if 1 MHz then 1024 us)

;

ldi mp,0b0000_0101

out TCCR0B,mp

;

select time between interrupts of approximately 10ms (10.24 ms)

;

ldi mp, 246

outTCNT0,mp

;

;	enables the interrupt in the event of a TCNT0 overflow

;

ldi mp,0b0000_0001

sts TIMSK0,mp

;

; initialises variables

;

ldi	var_check, T_CHECK ; variable that sets the voltage reading interval of the ;supply voltage

ldivar_check2	, T_CHECK2 	; variable setting the interval for reading the table

ldivar_check3	, 0b0000_0000			; control variable for anti-rebound

;

; initialises the variable flag_button which checks whether the button has been pressed

;

ldi flag_button, 0b0000_0000

;

; program the ADC to enable it without enabling the interrupt and select a prescaling factor of 4.

;

ldi 	mp,0b1000_0010

sts	ADCSRA,mp

;

; enables SREG-level interrupts

;

you are

;

;

;

main_loop:

;

; check whether a time equal to the voltage sampling interval has elapsed

;

cpivar_check	,0

;

; conditional jump to button if var_check is different from 0

;

brne	button

;

; initialise var_check and power control

;

ldi		var_check,T_CHECK

;

; 	conversion begins

;	I select the channel and voltage from the ADMUX

;

ldi		mp,0b0110_0000	; bit ADLAR = 1 for 8 bits, voltage at 5V, input A0

sts		ADMUX,mp

;

callDelay_1sec		           ; 5 ms delay to stabilise the voltage after the switch

;

; now start the conversion by setting bit 6 of ADCSRA (ADSC) to 1 without changing anything else in ADCSRA

;

lds		mp,ADCSRA

ldi		mp1,0b0100_0000

or		mp,mp1

sts		ADCSRA,mp

;

;

; wait until the conversion is ready by testing ADSC in ADCSRA: when the conversion is complete ADSC returns to 0

;

check_conv:

lds		mp,ADCSRA

ldi		mp1,0b0100_0000

and		mp,mp1

brne check_conv

;

;	read the converted value on 8 bits

;

lds		mp,ADCH

;

; if the read value is greater than or equal to the threshold, it jumps to next02

;

cpi		mp,TSH_BAT

brsh 	next01

;

; sets flags bit 0 to 1 without modifying the others

;

ldi		mp1,0b1111_1110

and		flags,mp1

;

next01:

cpi		mp,TSH_BAT

brlo 	next02

;

; sets flags bit 0 to 0 without modifying the others

;

ldi		mp1,0b0000_0001

or		flags,mp1

next02:

;

; check the flag to possibly switch on the LED connected to PB5

;

ldi		mp1,0b0000_0001			; prepares the mask for bit 0 of flags 0b0001_0000

and		mp1,flags			; isolates bit 0 of flags in mp1

cpi		mp1,0b0000_0000

brne 	next03

;

; These instructions write the value 1 bit 5 portB; the LED lights up

;

ldi mp,0b0010_0000

out PORTB,mp

;

next03:

;

; check the flag to possibly switch off the led

;

ldi		mp1,0b0000_0001			; prepares the mask for bit 0 of flags 0b0001_0000

and mp1,flags			; isolates bit 0 of flags in mp1

cpi		mp1,0b0000_0001

pulsating brne

;

; These instructions write the value 0 into the first line of port C; the LED turns off

;

ldi 	mp,0b0000_0000

out PORTB,mp

;

button:

;

; Firstly button status control: I insert a delay that allows me not to change state for 30 ms after ; that the button has been pressed,

Avoiding rebound artefacts due to its mechanical nature

;

sbicPIND	, PD7				; check pin 7 status, if PD7 is at 0 I switch to rjmp voltage

rjmp	skip

rjmp conv_voltage_temp

;

jump:

ldivar_check3	,30				; initialise var_check3 to 30

;

delay_30:						; 30 ms 	delay to avoid debounce

cpivar_check3	,0

brnedelay_30

ldi mp, 0b0000_0001

eor flag_button,mp

;

; check whether a time equal to the voltage sampling interval has elapsed

;

conv_tension_temp:

;

cpivar_check2	,0

;

; conditional jump to end_loop if var_check2 is different from 0

;

brne	end_loop

;

; if var_check2 is equal to 0

;

ldivar_check2,T_CHECK2			 ; initialisation var_check2

;

; check flag_button: if it is equal to 1 I read table_F otherwise I read table_C

;

andiflag_button	,1

brne	table_F

;

; prepares the passing of the initial table address to the subroutine as a local variable in the stack

;

table_C:

ldi	ZH,high(2*tableC)			  ; initialises ZH with the high part of the table address

ldi	ZL,low(2*tableC)			  ; initialises ZL with the lower part of the table address

;

pushZH			          ; puts ZH on the stack to pass it to the subroutine that will use it as a parameter

;

pushZL					  ; puts ZL on the stack to pass it to the subroutine that will use it as a parameter

;

rjmp	proof

;

;

table_F:

ldi		ZH,high(2*tableF)		  ; initialises ZH with the high part of the table address

ldi		ZL,low(2*tableF)		  ; initialises ZL with the low part of the table address

;

pushZH					  ; puts ZH on the stack to pass it to the subroutine that will use it as a parameter

;

pushZL					  ; puts ZL on the stack to pass it to the subroutine that will use it as a parameter

;

; call of the measure subroutine which samples the voltage and displays the result (note, the call stores the two bytes of the return address on the stack)

;

proof:

call 	measure

ret						  ; returns to the calling programme

;

;

brne	end_loop

;

end_loop:

;

nop						  ; nop = no operation

;

rjmp main_loop

;

; call up the measure subroutine which samples the voltage and displays the result

;

measure:

;

pop mp				; temporarily extracts the last byte of the return address placed on the stack by the call

pop		mp1				; temporarily pulls the first byte of the return address placed on the stack by the call

pop		ZL				; retrieves the low byte of the table address that was passed by the calling programme on the stack

pop		ZH				; retrieves the high byte of the table address that was passed by the calling programme on the stack

push		mp1				; restores the first byte of the return address placed on the stack by the call

push		mp				; restores the last byte of the return address placed on the stack by the call

;

;

;

ldi		mp,0b1100_0001	        ; set working voltage to 1.1 V, ADLAR=0, channel 1

sts		ADMUX,mp

call 	Delay_1sec

;

;

lds		mp,ADCSRA

ldi		mp1,0b010000	        ; start conversion

or		mp,mp1

sts		ADCSRA,mp

;

; wait until the conversion is ready by testing ADSC in ADCSRA: when the conversion is complete ADSC returns to 0

;

check_conv2:

lds		mp,ADCSRA

ldi		mp1,0b010000

and		mp,mp1

brne 	check_conv2

;

; reads the converted value on 10 bits ADCL and ADCH by reading ADCL first

;

; multiplication by 4 on a 10-bit register

;

lds		mp, ADCL

lds		mp1, ADCH

;

lsl		mp1

lsl		mp1				; I move the ADCL bits to the left, twice

;

lsl		mp

brcc 	carry				; switches to carry if carry=0, no carry

ori		mp1, 0b0000_0010		; if it is set (carry=1) I put the carry in the first bit of ADCH

;

carry:

lsl		mp

brcc 	display				; switches to display if carry=0, no carry

ori		mp1, 0b0000_0001 ; if set (carry=1) I put the carry in bit 0 of ADCH

;

; I show the converted values on the display

;

display:

;

add		ZL,mp				; Points to the table cell that corresponds to the value read by the ADC

adc		ZH,mp1

;

; prepares display

;

lpm		mp, Z+				; reads first value of the table row containing the complete integer value over 8 bits and increments Z

;

lpm		mp, Z+				; reads the second value of the table row containing the hundreds digit and increments Z

ori		mp,0b0111_0000	        ; leave one PD4 - PD7 and do not change PD0 - PD3

out		PORTD, mp			; writes out the hundreds digit

andi mp, 0b0110_1111	        ; prepares PD4 at low level (LE hundreds) without changing the hundreds digit

out		PORTD, mp			; enables LE hundreds

nop						; waiting instructions necessary to ensure strobe effectiveness

nop

nop

nop

nop

ori		mp,0b0111_0000	        ; puts at one PD4 - PD7 (raises LE)

out		PORTD, mp

;

;

lpm		mp, Z+				; reads the third value of the table row containing the tens digit and increments Z

;

ori		mp,0b0111_0000		 ; leaves PD4 - PD7 at 1 and does not change PD0 - PD3

out		PORTD, mp			; outputs the tens digit

andi mp, 0b0101_1111	            ; prepares PD5 at low level (LE tens) without changing the tens digit

out		PORTD, mp			; enables LE tens

nop						; waiting instructions necessary to ensure strobe effectiveness

nop

nop

nop

nop

ori		mp,0b0111_0000		 ; puts at one PD4 - PD7 (raises LE)

out		PORTD, mp

;

lpm		mp, Z+				; reads the fourth value of the table row containing the units digit and increments Z

;

ori		mp,0b0111_0000		; leaves PD4 - PD7 at 1 and does not change PD0 - PD3

out		PORTD, mp			; writes out the units digit

andi mp, 0b0011_1111		; prepares PD6 at low level (LE units) without changing the units digit

out		PORTD, mp			; enables LE units

nop						; waiting instructions necessary to ensure strobe effectiveness

nop

nop

nop

nop

ori		mp,0b0111_0000		 ; puts at one PD4 - PD7 (raises LE)

out		PORTD, mp

;

;	display writing finished

ret				        ; returns to the calling programme

;

;

;

Delay_1sec:

ldi dly1, 1

Delay1:

ldi dly2, 10

Delay2:

ldi dly3, 250

Delay3:

dec dly3

nop

brne Delay3

dec dly2

brne Delay2

dec dly1

brne Delay1

ret

; INTERRUPT HANDLERS FOLLOW

;

;

EXT_INT0:

networks							vector 2:		IRQ0 Handler

;

EXT_INT1:

networks							vector 3:		IRQ1 Handler

;

PCINTR0:

networks							vector 4:		PCINT0 Handler

;

PCINTR1:

networks							vector 5:		PCINT1 Handler

;

PCINTR2:

networks							vector 6:		PCINT2 Handler

;

WDT:

networks							vector 7:		Watchdog timer handler

;

TIM2_COMPA:

networks							vector 8:		Timer2 compare A handler

;

TIM2_COMPB:

networks							vector 9:		Timer2 compare B handler

;

TIM2_OVF:

networks							vector 10:		Timer2 Overflow Handler

;

TIM1_CAPT:

networks							vector 11:		Timer1 Capture Handler

;

TIM1_COMPA:

networks							vector 12:		Timer1 CompareA Handler

;

TIM1_COMPB:

networks							vector 13:		Timer1 CompareB Handler

;

TIM1_OVF:

networks							vector 14:		Timer1 Overflow Handler

;

TIM0_COMPA:

networks							vector 15:		Timer 0 CompareA handler

;

TIM0_COMPB:

networks							vector 16:		Timer 0 CompareB handler

;

;

TIM0_OVF:

;

; saves mp in the stack before using it

push	mp

; saves SREG in the stack

at		mp,SREG

push	mp

; re-initialises the counter timer count variable (10.24 ms)

ldi		mp,246

out		TCNT0,mp

; decreases the counting variables of the ADC timing

dec		var_check

dec		var_check2

; restores SREG

pop 	mp

out 	SREG,mp

; restore mp

pop 	mp

;

networks		                vector 17:		Timer0 Overflow Handler

;

;

;

SPI_STC:

networks		               vector 18:		SPI Transfer Complete Handler

;

USART_RXC:

networks		              vector 19:		USART RX Complete Handler

;

USART_UDRE:

networks		              vector 20:		USART UDR Empty Handler

;

USART_TXC:

networks		              vector 21:		USART TX Complete Handler

;

ADC_conv:

networks			 vector 22:		ADC Conversion Complete Handler

;

EE_RDY:

networks				vector 23:		EEPROM Ready Handler

;

ANA_COMP:

networks				vector 24:		Analog Comparator Handler

;

TWSI:

networks				vector 25:		Two-wire Serial Interface Handler

;

SPM_RDY:

networks				vector 26:		Store Program Memory Ready Handler